Register Map
Base Address 0x43C00000

Address (hex)	Register Name	Function
00	RESERVED	Reserved for future use.
04	RESERVED	Reserved for future use.
08	LEDS	Control LED's / bits 3-1 : Front Panel LED's
0C	FPGA_VERSION	FPGA firmware version (currently 121)
10	TOKEN	Controls token lines on detector board / bit 0 : token data / bit 1 : token clock
14	ASIC_SPI	Controls the SPI lines for Hermes and Scepter ASICS. / bit 0 : spi data / bit 1 : spi clock
18	CALPULSER	Hermes Calibration Pulser. Calibration pulse is synchronized to Read Request (RR) pulse / bit 0 : Single Calibration pulse / bit 16: Continuous Cal pulse / bits 31-24: Cal Pulse Rate if bit 16 is set. / 0 = cal pulse sent on every RR / 1 = cal pulse sent every other RR / 2 = cal pulse sent every third RR / 255 = cal pulse sent every 254th RR.
1C	SCEPTER_ENB	Connects to Scepter Enable Pin / 0 = scepter disabled, 1=scepter enabled
20	SCEPTER_RDOUT_ENB	Enables the Readout of Scepter ASIC / (see section xxx for full details of bits)
24	EVENT_TIME_RESET	Resets the event time counter. The event time counter is a 29 bit counter which is free running at 25MHz. The count value is latched and recorded with each event.
28	SIM_EVENT_SEL	For testing the readout when real events are not available, this will generate simulated events at a rate set by the SIM_EVENT_THROTTLE register / 0 = real events / 1 = simulated events
2C	SIM_EVENT_THROTTLE	Sets the event rate when the SIM_EVENT_SEL register is set to 1
30	ADC_SPI	Connects to the PD & TD ADC SPI port / bit 0 : sda / bit1 : sck / bit2 : csb
34	ADC_PWR_ENB	Puts the ADC in low power standy mode / 0 = active / 1 = power down
38	CALPULSER _DLY	Used in conjunction with CALPULSER register. Adds a delay between the RR pulse and the Calibration Pulse / bits 7-0 : CalDelay in 10ns steps.

60	EVENT_FIFO_DATA	Reads the Event Data
64	EVENT_FIFO_RDCNT	Number of events currently queued up to be read from the EVENT_FIFO_DATA register
68	EVENT_FIFO_CNTRL	Control Register for Event Data / bit 0 : fifo enable / bit 1 : testmode. Sends test count pattern for data / bit 2 : fifo reset

74	FEADC_TRIG	Triggers the Front End ADC to start a conversion / bit 2-0 : Selects ADC channel number / bit 3 : Initiate a conversion, must clear before starting a new conversion. / Readback is convert done
78	MAX1203_DATA	Readback of Front End ADC

80	DMA_CONTROL	Control Register for DMA operation / bit 0 : enable / bit 1 : testdata enable / bit 2 : fifo reset / bit 3 : test data counter clear / bit 4 : DMA mode. 0=continuous, 1=single BURSTLEN
84	DMA_STATUS	Status Register for DMA operation / bit 0 : dma active
88	DMA_BASEADDR	Base physical address for DMA transfers
8C	DMA_BURSTLEN	Number of 32 bit words to DMA
90	DMA_BUFLEN	DMA Buffer length (in lwords), DMA address will wrap back to the BASEADDR when it reaches the BUFLEN
94	DMA_CURADDR	Current Address to DMA write pointer
98	DMA_THROTTLE	When DMA is in testdata mode, sets the rate for DMA writes. 0=100Mwords/sec, 1=50Mwords/sec, ...
C0	DMA_IRQ_THROTTLE	Controls the IRQ rate for the DMA operations. Interrupt after every (n+1) 4kbyte transfers
C4	DMA_IRQ_ENABLE	Enable Interrupts after DMA operation
C8	DMA_IRQ_COUNT	Counter of number of interrupts sent

D0	FRAME_SOFTTRIG	1 = Start a new event frame. Readback reports if current frame is acitve
D4	FRAME_TIME	Sets the frame length. Resolution is 40ns. Interrupt is generated when the frame time expires. (IRQ # 61)
D8	FRAME_NUM	Sets the frame number on write. Read returns the current active frame number. Increments when frame completes
DC	FRAME_MODE	Sets the frame mode / 0 = debug mode. Frame length is set to infinite. / 1 = normal mode. Frame starts on FRAME_SOFTTRIG and completes after FRAME_TIME